FLUE / PROCESS GAS ANALYSER SWG 200-1

Continuous emission monitoring systems are designed to control and

monitor the emission level of combustion sources or industrial processes.

The standard following components can be measured and indicated:

O2     Oxygen			 0 – 25,0 % electrochemical cell

CO  Carbon monoxide 	 0 – 4.000 ppm NDIR bench

CO2  Carbon dioxide		 0 – 20,0 % NDIR bench

CH4  Methane 		 0 – 10.000 ppm NDIR bench

SHORT DESCRIPTION

• The continuous emission analysis is based on infrared bench (NDIR) and long life electrochemical cell.

• The SWG 200-1 will be installed nearly to the gas sampling points, so that there is only

a short distance between the analyzer and the gas sampling point.

This will decrease the response time of analyzer.

The standard length of the non-heated PTFE gas sample line is 10 m.

• The SWG 200-1 is equipped with a complete sample conditioning system.

• The SWG 200-1 is equipped with a standardised digital data interface and analogue

outputs as 0/4-20mA current loop. The measured data will be provided at these outputs.

• Display for analysis data and service data indication are direct in front of the SWG 200-1.

• The SWG 200-1 is designed for continuous measurement. The automatically zero calibration

is occurring at least once a day according to the free user software settings.

• By means of inlets at the bottom of the enclosure, the installation of the power supply 230V / 50Hz,

( 115V / 50-60Hz on request) the gas sampling line, the condense outlet and the data lines are carried out.

• The place of installation SWG 200-1 is in a dry, weatherproof ambient with temperature between

+5°C and +40°C.

HARDWARE

The complete hardware is mounted into a steel cabinet

(weight approx. 30kg, protection IP 43, with 4 fastening loops for wall mounting).

- Lockable glass door

- Local operation keyboard and back lighted LCD display for indication of measured data.

- Complete flue gas conditioning system with:

- Gas sampling pump

- Sample gas cooler

- Automatic condensate draining

- Filter elements for flue gas and ambient air

- Flow meter with needle valve

- Solenoid valve for auto-zero

- µ-processor electronic board

- Sensors and amplifier electronics, infrared bench

- RS485 interface for long distance data transfer

- 8x analog outputs as 4-20mA current loop

SOFTWARE

Standard software for all aforementioned functions with self-check of all internal functions:

• System diagnosis

• Peltier cooler automatic control

• Sensor functions

• Condensate monitoring

• Automatic and standby operation

• Digital data transmission, including PC visualization software

Analyzer specifications

| Measuring principle: | non dispersive infrared, single light source with electronic modulation, single beam system, multigas detector for CO, CO2 and CH4, and/or long life electrochemical cells for O2 , CO, NO and NO2 |
| --- | --- |
| Measurable components and measuring ranges : | CO: 0 – 4.000 ppm CO2: 0 – 20 / 40 % CH4: 0 – 10.000 ppm O2 : 0 – 25,0 % Other measuring ranges on request |
| Resolution: | 1 ppm respectively 0,1% |
| Detection limit: | 20 ppm respectively 0,1% |
| Repeatability: | +/- 3% FS |
| Linearity: | +/- 3% FS |
| Long time zero drift: | none, with automatic auto-zero |
| Span drift: | ess than +/- 5% FS over 1 week |
| Temperature effect on zero | none, with automatic auto-zero |
| Temperature effect on span: | none, internal temperature compensation |
| Response time: | less than 30 seconds for 90% FS response, from analyzer inlet port |
| Warm-up time: | minimum 1/2 hour |
| Gas sample conditioner: | flow rate: 0,5 l/min amb. temp.: 0 - 40°C moisture: less than +5°C dew point dust: less than 1um inlet pressure: 700 mbar….1200mbar absolute inlet/outlet ports: 6 mm fitting |
| Calibration: | by software, calibration gases for every gas required clean and dry ambient air for auto-zero purging |
| Sample gas flow: | flow meter with needle valve, stainless steel, flow monitoring and alarm |
| Power supply: | 230 Vac / 50 Hz, 100 W ( other on request) |
| Electronics: | microprocessor, graphic display with backlight, membrane keypad |
| Analog output: | up to 8 channel 4 - 20 mA |
| Digital output: | RS485 for long distance data transfer |
| Digital input: | 3x potential free voltage, 12V to 24V ac or dc |
| Electrical connections: | rear panel 19” rack feed-through terminal blocks |
| Ambient temperature: | +5°C ... +40°C |
| Ambient humidity: | up to 95% RH , non condensing |
| Enclosure: | Steel, high protection finish for IP43 |
| Dimensions & weight: | 600 x 220 x 4200 mm ( W x H x D ), 24 kg |

Gas sampling probe LD, with in-situ filter

Gas sampling probe for applications with low dust rate, with:

• Stainless steel probe tube length ∅ 22 mm x 250/500/750/1000 mm,

or longer up to max. 2000 mm, for flue gas temperatures up to 500°C

• Inconel steel probe tube length ∅ 22 mm x 500/750/1000 mm,

or longer up to max. 2000 mm, for flue gas temperatures up to 900°C

• Stainless steel 1.4571 (SS316Ti) flange DN65 PN6

• In-situ sintered metal filter 3 um, screwable,

pipe connection for Teflon gas sample line DN6/4

with Teflon gas sample line DN6/4, 10 m

Option:

• Heated gas sample line with:

o exchangeable Teflon hose (PTFE), DN6/4

o 230Vac, 100 W/m

o length 5 m (for other length see price/meter)

Gas sampling probe HD, with heated filter

Gas sampling probe for applications with high dust rate, with:

• Stainless steel probe tube ∅ 22 mm x length 250/500/750/1000 mm,

or longer up to max. 2000 mm, for flue gas temperature up to 600°C

• Inconel steel probe tube ∅ 22 mm x length 500/750/1000 mm,

or longer up to max. 2000 mm, for flue gas temperature up to 1100°C

• Ceramic probe tube ∅ 22 mm x length 750/1000 mm,

or longer up to max. 2000 mm, for flue gas temperature up to 1700°C

• Stainless steel flange DN65 PN6

• Heated ceramic filter to approx. 150 °C, easy access for maintenance

• Automatic back purging with oil and water free compressed air, connection 6...10 bar,

• Pipe connection for heated sample line

• Protection cover against the hot surface of the filter

Option:

• Heated gas sample line with:

exchangeable Teflon hose (PTFE), DN6/4

230Vac, 100 W/m

length 5 m (for other length see price/meter)